## Slide 1: Development of a Segmented Planar Germanium Imaging Detector
University of Liverpool, STFC Daresbury Laboratory,
e2v, Ortec
Germanium Development Proposal

## Slide 2: The Objective of the Project
The objective of the project is “to develop the capability in the UK to realise a compact highly segmented planar germanium detector system capable of gamma-ray imaging”.
The project aims to get one and then two detector elements working in a custom designed cryostat.
To characterise the performance of the detector system.
To optimise performance for applications that focus on the use of Compton Imaging techniques.
A proposal will then be developed in collaboration with e2v for future exploitation of the capability for SPECT imaging, Homeland Security and DESPEC at NuSTAR.
Germanium Development Proposal

## Slide 3: The System Requirements
A high performance Gamma-ray Imaging system requires:
Detector performance
  Excellent energy resolution
  Good time resolution
  Good position resolution
A compact geometry
  Maximise system efficiency
  Minimise dead material
Electronic collimation
  Maximises system efficiency
PRD Germanium Development Proposal

## Slide 4: HPGe detector manufacturing
PRD Germanium Development Proposal
<number>
Poly Crystalline
Zone
Refiner
Crystal
Puller
Ge Single
Crystal
Boron
Implant
Wet
Lab
Lithium
Diffusion
Mechanical
Preparation
Load in
Cryostat
IV
Test
Spectroscopy Test
Mechanical
& electronic
Ship
Thermal
Cycle
Good?
Good?
Good?
Good?
Courtesy Ortec

## Slide 5: Germanium Detector Development
Steps required to deliver a segmented germanium detector:
A high quality supplier of zone refined germanium
Cut and shape the crystal with specific axis orientation
Contact the crystal with lithographic masking
Cryostat design and build
Preamplifier selection and integration
System Integration
System Characterisation
Germanium Development Proposal

## Slide 6: The Track Record
PRD Germanium Development Proposal

## Slide 7
Double Sided HPGe Strip Detectors
60mm x 60mm x 20mm active area
7mm x 20mm guard ring
12 x 12 orthogonal strips
- 5mm pitch
- 5mm x 5mm x 20mm voxels
1mm Aluminium entrance window
Thin contact technology
Fast charge sensitive preamplifiers
Energy resolution:
1.5 keV@122 keV & 3.25keV FWHM at 511keV
Intrinsic photopeak efficiency – 19% at 511keV
SmartPET detectors
PRD Germanium Development Proposal
H. Boston et al., A579(2007) 104-107
Notes: AC – 0.3um, 180um separation, DC – 50um, 300um separation
<number>

## Slide 8
Three 22Na point sources have been imaged with the SmartPET system
60mm
60mm
From MLEM reconstruction the point sources display FHWM of ~1.4mm
Over 60% of events processed
Point Source Imaging
PRD Germanium Development Proposal
R. Cooper, H. Boston, NIMA 579 (2007) 313-317

## Slide 9: Imaging Progress : Compton Camera
PRD Germanium Development Proposal
Cone beam reconstruction with 10 iterations.
~8mm image resolution x-y.
152Eu point source imaging.
30 keV gate on 1408 keV.
30mm detector separation with 1.6mm position resolution.
Single interactions in each detector.
6 cm source to crystal
3 cm crystal to crystal
H. Boston et al., NIMA580 (2007) 929-933

## Slide 10: Compton Imaging with HPGe
PRD Germanium Development Proposal
30mm & 50mm separation between scatterer & analyser.
1.6cm separation between points
FWHM ~ 8mm

## Slide 11: The Potential Applications
Examples of the application of such detectors
Germanium Development Proposal

## Slide 12: The Potential Applications: SPECT
Motivation: Factor of 100 increase in system sensitivity.
The development of a SPECT demonstrator. The system will be capable of imaging 99mTc (141 keV) and other low energy SPECT isotopes.
Utilise Compton Image reconstruction techniques, to electronically collimate the gamma-rays.
Follow on the from the successful SmartPET project
However, will require more than one germanium crystal in a single cryostat.
Work with UK and international industry.
Germanium Development Proposal

## Slide 13: The Potential Applications: Homeland Security
Motivation: Multi-nuclide gamma-ray imaging
  High resolution spectroscopic capability
  (60keV – 10 MeV)
  High sensitivity (with large field of view)
  Excellent imaging capability (<8mm FWHM for reconstructed point source)
For Environmental Assay
For (n,n’), (n,) illicit substance detection
Nuclear waste management and decommissioning
  The identification, location and quantitative activity assessment of radioactive sources in 3-dimensional structures
  Imaging in mixed radiation fields or high field imaging, e.g., integrity of containers, identification of contents using planar or tomographic techniques.
Germanium Development Proposal

## Slide 14: The Potential Applications: DESPEC
Germanium Development Proposal

## Slide 15: Advances in DESPEC gamma detection
Background Rejection
Imaging determines the location of the gamma-ray photon. This coupled with a signal from AIDA in the focal plane of the SFRS reduces background by factor of 10, keeping 80% of important events. Improves PT by factor 100. [1]
Prompt Flash
High granularity allows detector system to recover from prompt gamma flash. Leads to access to shorter lived states and new physics.
Geometry
Designed for AIDA, giving excellent solid angle coverage.
Germanium Development Proposal
[1] S. Tashenov, J. Gerl, NIMA 586 (2008) 224–228

## Slide 16: Status & Milestones
Germanium Development Proposal

## Slide 17: Project Milestones
October 2008 – Project start
November 2008 – System specification complete (WP2)
January 2009 – Detector mounts complete (WP3)
May 2009 – Integration of detector 1 with mount complete (WP3)
May 2009 – Cryostat 1 complete (WP4)
July 2009 – Preamplifier integration complete (WP5)
Oct 2009 – System integration complete with crystal 1 (WP6)
Dec 2009 – One detector system performance characterised (WP7)
Feb 2010 – Integration of detector 2 with mount complete (WP3)
May 2010 – System integration with complete with crystals 1 + 2 (WP6)
Sept 2010 – Two detector system performance characterised (WP7)
Germanium Development Proposal
Funding decision due summer 08

## Slide 18: Development of a Segmented Planar Germanium Imaging Detector
University of Liverpool, STFC Daresbury Laboratory, e2v, Ortec
Germanium Development Proposal