## Slide 1: AGATA Detector Work Performed at Saclay in 2008 - lessons learned -
Julien Pancin
Joa Ljungvall
Alexandre Obertelli
Marc Karolak
Andreas Görgen

## Slide 2
03.09.2007
Saclay test cryostat (TestC2) sent to CTT for repair because of vacuum leak
27.01.2008
received test cryostat TestC6 (owner Padova) from GSI with S003 mounted
received A005, B005, C005 (owner Sweden) from Cologne for CAT
plan: test S003 to validate / demonstrate functioning of the system, in particular
good functioning of the cryostat, then test the 3 asymmetric detectors
19.02.2008
good energy resolution
with analogue system
for all segments and core
“plug and play”

## Slide 3
25.02.2008
crosstalk measurement with digital TNT2 system
no unusual crosstalk observed
conclusion:
S003, TestC6, and DAQ system in good shape
ready for CAT of A005, B005, C005

## Slide 4
10.03.2008
A005 mounted in TestC6 for CAT
observations:
far too much noise
core resolution 3.2 keV at 1.3 MeV
several segments at ~2 keV at 60 keV
oscillations of various frequency
base line unstable
suspicion:
grounding problem
19.03.2008
warm up, all mass connections checked, everything seems okay
26.03.2008
cooling down again
28.03.2008
lost core FET while raising HV (at 2600 V)
(first time this happened, unclear why)

## Slide 5
09.04.2008
cold core preamp repaired, A005 cold again
HUGE crosstalk between core and all segments (on the scope)
reduced when PT100 read-out disconnected
14.04.2008
A005 warm and open
cross talk caused by contact between shielding of PT100 cable and
shielding of cold core preamp
decision to test cryostat again with S003
25.04.2008
S003: excellent resolution of core and all channels

## Slide 6
28.04.2008
A005 mounted again
06.05.2008
detector cold, again too much noise, unstable base line
spontaneous oscillations observed
possible to suppress oscillations, but not for long
16.05.2008
stable operation not possible, warm up

## Slide 7
19.05.2008
removed all cold segment preamps, all segment signals connected to ground
test core only
23.05.2008
detector cold
unstable base line, far to much noise, susceptible to oscillations
core resolution ~4 keV (segments still grounded)
questions
are asymmetric detectors more sensitive than symmetric ones ?
are there any other differences beside shape of the crystal ?
is A005 really bad ?
is the cryostat tuned to a given detector and only to that one ?
suspicion
starting to doubt grounding scheme in cryostat
warm up

## Slide 8
26.05.2008
changed ground connection of cold preamp board to cold finger
29.05.2008
A005 cold again
oscillations and noise unchanged

## Slide 9
02.06.2008
decided that A005 should be tested in Cologne with different cryostat
mounted B005 in TestC6
05.06.2008
B005 cold
system oscillating just like A005
questions
are asymmetric detectors more sensitive than symmetric ones ?
are there any other differences beside shape of the crystal ?
are both A005 and B005 really bad ?
is the cryostat tuned to S003 and only to that one ?

## Slide 10
09.06.2008
investigated cold preamp board
internal ground layer on front and back side isolated
ground of individual blocks isolated from internal layers and from each other
ground of 6 blocks connected to each other and to front layer via external cable
ground layer on the back floating !
board mounted on S003 has no internal layers
finally we found a difference !

## Slide 11
the boards have evolved in time
there are at least 5 different versions: with or w/o internal ground layers
isolated or connected
not only that: grounding scheme in the various test cryostats differs also
(triple cryostats are different yet again)
consequence:
detectors are not easily interchangeable and need “bricolage” for every individual case
need to establish standard
must be possible to mount any detector in any cryostat without changing the grounding

## Slide 12
09.06.2008
connected front and back side via external cables on B005
12.06.2008
B005 cold
noise level for all segments normal
no oscillations observed
discussions with Cologne, more detectors with floating ground layer found
these were causing oscillations in the triple cryostat
core resolution still poor
noise band of the core reduced when changing shaping time from 6 to 3 µs
indication of leakage current between core and a segment ?
shaping time 6 µs
shaping time 3 µs

## Slide 13
TestC6 is not up to date, has no pulser functionality, and needs “debugging”
TestC2 (owner Saclay) has been repaired, equipped with pulser, and is state of the art
dewar repaired
capton signal cables replaced
by individual cables
2 additional ground connections
from cold FET directly to warm
preamp ground
equipped with pulser
new plan:
go to Cologne
test rebuilt Saclay TestC2 with accepted detector (A002)
then perform CAT of A005 in TestC2 at Cologne
solve problems of TestC6
17.06.2008
TestC6, A005, and B005 transported to Cologne
visits of J.Pancin, M.Karolak, and A.G. several days each

## Slide 14
18.06.2008
mounted A002 in TestC2 at Cologne
TestC2 had been tested with dummy before
changed grounding scheme (with H.-G. Thomas)
tested with dummy – no oscillations
(H.-G. Thomas)
A002 shows oscillations (A.G. 19.6.)
oscillations confirmed with dummy
(J. Pancin 23.6.)
no oscillations with dummy
A002 is working well (M. Karolak 25.6.)
core slightly above specs
influence of core pulser – why?
3.7.2008 A005 mounted (Cologne team)
8.7.2008 testing (J. Pancin)
problems with segments
warming up

## Slide 15: General remarks (1)
Good collaboration between Cologne and Saclay
We had very useful EVO meetings – should take place regularly
electronic log book very useful – but we need to use it

## Slide 16: General remarks (2)
Microphonic behavior has not been observed after gluing
problem seems solved
need to establish standards
need documentation
planning:
perform CAT of A005 in TestC2 in Cologne
in collaboration between Cologne and Saclay
move TestC2 (and eventually TestC6) back to Saclay
perform CAT of B005 and C005 at Saclay
prepare S003 in TestC6 for scanning in Orsay