## Slide 1: Detectors and Preamplifiers
Detector deliveries
Customer Acceptance Tests
Available detectors
AGATA triple detectors
Scanning program
P. Reiter IKP University Cologne AGATA week 8-11 July 2008 Uppsala
Notes: <number>

## Slide 2
Detector | Delivery | Owner | Status
A001* | 11/2005 | GANIL | rejected – hole trapping
A001 | 06/2006 | GANIL | accepted
A002* | 04/2006 | TU Munich | rejected – leakage currents
A002 | 06/2007 | TU Munich | accepted
A003* | 11/2006 | Liverpool | rejected – leakage currents
A003 | 10/2007 | Liverpool | accepted
A004 | 09/2007 | Ankara | rejected – segments out of specification
A004* | 07/2008 | Ankara | CAT at IKP, Cologne
A005 | 12/2007 | Stockholm | CAT at CEA, Saclay & IKP
B001* | 11/2005 | Padova | rejected – segments out of specification
B001 | 06/2007 | Padova | rejected – leakage currents
B002 | 08/2006 | GANIL | accepted
B003 | 06/2007 | Liverpool | accepted
B004 | 12/2007 | Ankara | rejected – missing segment in capsule
B004* | 06/2008 | Ankara | CAT at IKP, Cologne
B005 | 12/2007 | Stockholm | CAT at CEA, Saclay
C001 | 12/2005 | Padova | accepted
C002 | 07/2006 | GANIL | accepted
C003 | 03/2007 | Liverpool | accepted, broken, leakage currents
C004 | 09/2007 | Ankara | accepted
C005 | 12/2007 | Stockholm | CAT at CEA, Saclay
Notes: <number>

## Slide 3
Detector | Delivery | Owner | Status
A001 | 06/2006 | GANIL | Accepted, available
A002 | 06/2007 | TU Munich | Accepted, available
A003 | 10/2007 | Liverpool | Accepted, available
B002 | 08/2006 | GANIL | Accepted, available
B003 | 06/2007 | Liverpool | Accepted, available
C001 | 12/2005 | Padova | Accepted, available
C002 | 07/2006 | GANIL | Accepted, available
C004 | 09/2007 | Ankara | accepted
Notes: <number>

## Slide 4
Recent detector tests and future deliveries
Turkish detectors
B004 (repaired) delivered to IKP, Cologne, test started
A004 (repaired) delivered July, 1st 2008
shipment of A004, B004, C004 to Turkey autumn 2008
Swedish detectors
A005, B005, C005 tests ongoing at CEA
- test results of A005 to be confirmed at IKP, Cologne
- test of B005 started
Liverpool detector C003: leakage current, out of warranty
Italian detectors
A006 announced delivery July (week 29) 2008
B006 announced delivery July (week 31) 2008
B001 announced delivery July 2008
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 5
Summary available detector
8 detectors accepted (A1,A2,A3 B2,B3 C1,C2,C4)
1 detector to be repaired at Canberra (B1)
5 detectors ongoing test (A4, A5, B4, B5, C5)
1 detector broken (C3)
----------------------------------------------------------------------------------------
3 (Sweden A005, B005, C005) ongoing CAT at Saclay
2 (Turkey A004, B004) ongoing CAT at Cologne
1 (Italy C001) scanning at Liverpool until July
1 (Italy B001) not accepted after CAT, not repaired
1 (UK C003) is broken after CAT, needs repair
3 (Turkey A004, B004, C004) are shipped to Turkey fall 2008
-----------------------------------------------------------------------------------------
At the moment 6 detectors are available for AGATA triple clusters
A001, A002, A003
B002, B003
C002
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 6
AGATA triple cluster detectors
First AGATA triple *
A001 :
B002 :
C002 :
*Goes to Legnaro
July 2008
Second AGATA triple
A003 :
B003 :
C004 : *
*Goes to Turkey in fall 2008,
Replaced by C005
Third AGATA triple
A002 :
B005 or B001: *
C001: **
*B005 after test or B001 after delivery & test
** C001 after scanning
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 7
Scanning program
Liverpool
C001 *
A005
* Goes to IKP,
Assembly of second
Triple cluster by CTT
July 2008
Orsay
S002 *
B005 :
* Commissioning, confirm
Liverpool results
See talks by
A. Boston
C. Unsworth
A. Korichi
F. Crespi
Thursday, 14:00-15:40
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 8
Future detector tests and triple assembly
Detectors for period fall and winter 2008/09
13 detectors may be available
A1, A2, A3, A5, A6
B1, B2, B3, B5, B6
C1, C2, C5
A6, B1, B6 delivery July 2008
redelivery of turkish detectors A4, B4, C4 ?
January - June 2009 italian detectors C6, A7 ?
detector repair of C3 ?
summer 2008 order of german detectors A8, B7, C7 ?
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 9
Talks by
A. Wiens, IKP Cologne
A. Goergen, CEA Saclay
D. Quirion, Canberra
B. Bruyneel, IKP Cologne
G. Pascovici, IKP Cologne
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 10
1st Tripel Cluster Energies: Single vs Triple cryostat
A001 – FWHM [keV]
B002 – FWHM [keV]
C002 – FWHM [keV]
Resolution 60keV line
Resolution 1.33MeV line
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 11
2nd Tripel Cluster Energies: Single vs Triple cryostat
A003 FWHM [keV]
C004 FWHM [keV] B003 FWHM [keV] A003 FWHM [keV]
Resolution:
60keV line
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 12
Tripel Cluster Energy resolution:
Single vs Triple cryostat
Excellent energy resolution in two AGATA triple clusters
Mean values segments @ 60 keV :
especially at low energies (electronic noise):
- glued detectors  less microphonics
- HEKO  AGATA preamp
- electronic assembly (grounding, cabling)
First AGATA triple
A001 : 1011 +/- 53 eV
B002 : 1039 +/- 70 eV
C002 : 965 +/- 63 eV
Second AGATA triple
A003 : 1052 +/- 65 eV
B003 : 995 +/- 73 eV
C004 : 1148 +/- 104 eV *
* high frequency oscillation to be removed
Notes: <number>
Improved energy resolution especially at low energies (electronics dominated)
- glued detector less microphonics
- final preamp
electronic assembly (grounding, cabeling)
Message : better than in single cryostat – specification easily met

## Slide 13: Cross talk measurements
A2..A6
B1
F1
Core to seg Xtalk
= understood!
Segment labeling:
Sectors: A...F
Rings: 1...6
X
Bart Bruyneel, Andreas Wiens
Notes: <number>

## Slide 14
Smallest capacities in ring 2
A2
B2
C2
D2
E2
F2
Notes: <number>

## Slide 15
Cross talk in first AGATA Triple Cryostat
A001
B002
C002
Core B
Core C
Core A
Core C
Core A
Core B
Sector A
Sector F
Sector A
Sector F
Sector A
Sector F
A001
B002
C002
Notes: <number>

## Slide 16
AGATA Triple Clusters: Summary & outlook
Two AGATA Triple Cluster Detectors operational:
- first & second triple: cryostat, mechanics,
vacuum, thermal properties of preamps o.k.
- seg-preamps, core preamps & pulser
with high dynamic range operational
- very good energy resolution at low energies
- no cross talk between detectors observed
in 1st ATC between detectors
- commissioning of 2nd ATC ongoing
Next steps:
- transport of 1st ATC to Legnaro, first tests
with digital electronics at Legnaro July 2008
- finish commissioning of 2nd ATC,
cross talk tests
- assembly of 3rd ATC constrained by
available detectors
ATC1
ATC2
ATC3
Notes: <number>