## Slide 1: AGATA week Uppsala, July 2008
Singles and Coincidence
Scanning of C001 in Liverpool
Notes: <number>

## Slide 2: Energy Resolution Measurements
FWHM measured in Liverpool (top) at 60 keV (blue) and 1332 keV (red).
C001 coupled to ORTEC 671 spectroscopy amplifier and ORTEC ADC.
Count rate maintained below 2 kHz on core.
These results are as good or better than those found in Cologne (bottom)
Segment F4 missing from data due to broken connection between FET and preamp.
Notes: <number>

## Slide 3: Efficiency Measurements
Absolute efficiencies were measured using 241Am and 152Eu sources placed 25cm from the front face of the crystal.
Results for C001 are consistent with those seen for S002.
εABS = 0.30% @ 120 keV
εABS = 0.08% @ 1408 keV
Notes: <number>

## Slide 4: Scanning table Setup
C001 mounted on coincidence scanning assembly.
Collimated 137Cs source moved in x-y plane on Parker scanning table.
Interaction position of 90˚ Compton scatter events identified through coincidence between C001 and scintillators mounted at 6 depths.
When system is triggered all 37 signals are recorded in 128 12.5ns samples using 14 GRT4 cards.
Notes: <number>

## Slide 5: Time Alignment
Time alignment has been the limiting factor in analysis of scan data from symmetric prototype detectors.
3 AGATA signals per 4 channel GRT card.
1 channel per card fed with triangular pulser.
After interpolation linear fit to both sides of pulse will identify peak sample.
All signals then aligned to this point.
Notes: <number>

## Slide 6: Segment Gated Intensity
Number of counts at each scanning table position when 662 keV is deposited in a single segment during singles scanning.
Notes: <number>

## Slide 7: Crystal Alignment
Two methods applied to intensity map data to obtain crystal tilt.
Method 1 finds the minimum bounding circle of points near centre which fall below threshold number of counts.
Method 2 uses a cross correlation between the intensity map for each ring and an image of a single circle.
0.38 mm between centres of rings 2 – 6.
0.88 mm between centres of ring 1 and 2.
Notes: <number>

## Slide 8: MWD Energy Spectra
Notes: <number>

## Slide 9: Core Risetimes (T30)
t30 – t10 for fold 1 photopeak events.
Fast axis (100) not aligned with any segment centre.
Notes: <number>

## Slide 10: Core Risetimes (T90)
t90 – t10 for fold 1 photopeak events.
Weak field evident at corners of ring 1.
Notes: <number>

## Slide 11: Segment Risetimes (T30)
t30 – t10 for events when 662 keV deposited in one segment, for rings 1 to 6.
Notes: <number>

## Slide 12: Segment Risetimes (T90)
Long T90 at corners of ring 1 due to weak field.
Notes: <number>

## Slide 13: Image Charge Asymmetry
ICA = [1000 x (Aclock – Aanti) / (Aclock + Aanti)] + 1000
Reduced sensitivity in mid radius regions of segments.
Notes: <number>

## Slide 14: Crosstalk Measurements
Analogue crosstalk measurements unreliable due to problems with spectroscopy amplifiers and ADC’s.
Digital crosstalk measurements to commence shortly after return to Liverpool using 60Co source and baseline offset measurements from GRT4 cards.
Notes: <number>

## Slide 15: Thank You
Collaborators
Mathew Dimmock
Steve Barrett
Andy Boston
Helen Boston
Ren Cooper
John Creswell
Notes: <number>