## Slide 1: Status of the AGATA scanning table using TIGRESS cards at Orsay
Agata week Uppsala July 2008
HA Thi Mai Hoa

## Slide 2
 Test a NEW electronics: TIGRESS-Ethernet link
 12 November: first TIGRESS card for test.
July 2007: received Exogam detector --> DAQ: TNT2 + Narval
Option Energy + Trace: instable DAQ due to USB links (high rate data flow)
Uncorrelated data due to the fact that the trigger number for trace and energy are different.
Rise time distribution T90
x (mm/10)
y (mm/10)

## Slide 3
50 electronics channels
(5 TIG10) + 1 collector card(TIG C)
VME crate
Flash ADC 14bits 100MHz
June 08: commissioning of this electronics started with standard EG Ge detector
(with help from J.P Martens_Montreal University)
Test with 1 TIGRESS card: OK
May 08 we were equiped with:

## Slide 4: TIGRESS card tests: signal and resolution (1)
Pulser @ different rates: 10Hz, 50Hz, 100Hz, 1kHz, 10kHz --> ok
EG Ge + NaI detector : test with radioactive sources.
NaI TOHR
Pulser
EG Ge det
t(x10ns)
t(x10ns)
t(x10ns)
Amplitude
Amplitude
Amplitude
( NaI TOHR is used to detect the scattered γ)

## Slide 5: TIGRESS card tests: signal and resolution (2)
FWHM with TIG10: 1.63 keV @ 121 keV
2.22 keV @ 662 keV
2.80 keV @1408 keV
(corresponds to the measured resolution before coming to Orsay)
--> The card gives the performances as expected in term of resolution.
Spectra EG Ge with 152Eu and 137Cs sources
662 keV
channel
Num of events
channel
Num of events

## Slide 6: Single scan standard EG detector (energy+trace)
T30
Photo peak Intensity
T60
T90
2mm step, 10s/position, 100 points/trace.
Energy distribution
  (gate on 662 keV photo peak)
Rise time distribution
  (T90, T60, T30- 662 keV gated)
x (mm/10)
y(mm/10)

## Slide 7: Coincidence measurements (1)
1. Coinc Ge – 1 TOHR module (22Na)
511 keV
Ge
NaI
Analysis off line using timestamp and on line using a trigger: same result.
Ge
NaI
Coinc matrix Ge - 1 TOHR module
Ge coincidence spectra
NaI coincidence spectra
Schematic of coinc set up Ge-1 TOHR module
channel
channel
Num of events
Num of events

## Slide 8: Coincidence measurements (2)
2. Coinc Ge - 6 TOHR (137Cs)
EG det
137Cs source: 477 Mbq (~ 1.8 kHz after collimator)
6 TOHR modules: - look at the same focal point
- 2 groups: 3NaI are positioned at 79.8°
3NaI are positioned at 100.2°
Compton scattering formula:
So, @ 79.8°: Energy deposited in Ge = 340keV and in NaI = 322keV
@100.2°: Energy deposited in Ge = 401keV and in NaI = 261keV
Resolution of NaI: 12%-17%
Schematic of coinc set up Ge – 6 TOHR modules

## Slide 9
Energy range in NaI: 240-360 keV
in Ge : 300-420 keV
This energy range can be explained by:
Energy/2 in Ge(keV)
Coincidence matrix Ge - 6TOHR
Energy/2 in Ge (keV)
Energy/5 in NaI (keV)
Energy/5 in NaI (keV)
Coinc matrix Ge - TOHR@ 80°
Coinc matrix Ge - TOHR@ 100°
Resolution of NaI
Angular coverage

## Slide 10
The number of single interaction Compton scatter events at a given position :
  From GEANT4 simulations for AGATA detector
400/h @ front face
100/h @ front third
10/h @ the back of the detector
  From measurement with standard EG detector
358/h @ front face
177/h @ 2cm from the front face
☺
Pulse shape of coincidence event at differents radius (given z)
t(10ns)
t(10ns)
Amplitude
Amplitude
30 14 10
mm
All signals @ r=14mm

## Slide 11
SCANNING TABLE @ ORSAY IS READY NOW
Stability of DAQ: ok
Energy resolution: ok
Counting rate: ok
Intensity and rise time distributions: ok
Coincidence measurement: similar to simulation
To conclude: with the new electronics TIGRESS