## Slide 1: Analysis of S002 scan data
Bart Bruyneel and Benedikt Birkenbach – IKP Cologne
At the 7th AGATA week, Uppsala July. 2008
S003

## Slide 2: Procedure
Scan data:
S003 scan @ liverpool
Averaged by J. Ljungvall
Simulation (~20 parameters):
Positions (offsets)
Responses (Preamp + Xtalk)
Mobility (e-, h+ and lattice)
Fields (Space charge)
Comparison:
FOM (Chi Square)
Repeat till parameters optimal

## Slide 3: Chi Square evaluation
> Almost 2x worse in Chi Square <
??? Noise, derivative crosstalk, ... ???

## Slide 4: Parameters in detail : a) Offsets
Position of crystal:
X,Y position known from singles scan,
Z position to be determined relative to collimators
Optimum S002: Zoffset = 3,86 mm
S003: Zoffset = 5,24 mm
Variation of lattice orientation:
S002 : <100> = -14,9º
S003 : <100> = 43.7º
S002 : <100> = -14,9º
S003 : <100> = 43.7º
A
F
B
E
D
C

## Slide 5: Parameters in detail : b) Mobilities
Electron mobility equal
Hole mobility unequal

## Slide 6: Parameters in detail : c) Space Charge
Zoom:
Rather insensitive in certain regions.
Region > 0 is not physical
Region < -2 is not likely
<-2,0
-1,0
-0,1
-1,0
Core @ 5000V
Unphysical
Space Charge [1010/cm3]
cylinder symmetry
Linear in r and z :
 4 point fit
S002:
S003:
-3,0
-2,9
1,3
1,1
Core @ 4000V
Unphysical
Unrealistic model for S003 !!!

## Slide 7: Option: Space charge from capacity measurements
50% variation
Late depletion
“Capacitance-voltage profiling„
e.g. Review article: Semicond. Sci. Technol. 1 (1986) 7-27